Bio-Plex Assay Notes

Pam Kreeger

06/08/06

Although this is not a complete list, here are some common things to watch out for.  For more tips, check the TroubleShooting Section of the product kit.

- Be careful with filter plates used for incubations

  - Pipet onto side and be careful not to puncture bottom

  - Calibrate vacuum manifold to 1-2.5 (in Hg) before EVERY wash step with a standard 96 well plate.  Too high a vacuum can cause the beads to get stuck in filter or damage filter.

  - Use fresh sealing tape after each step to completely seal plate and prevent cross-contamination of wells.

  - Unused wells of filter plates can be reused- simply keep the unused area sealed during 1st assay procedure (temporarily remove sealer when putting into instrument!), and mark off used wells when done.

  - If you need a round bottom plate, the needle height needs to be adjusted – talk to the Czar to get help for this.

- General Notes

  - Be sure to protect plate from light and keep sealed when shaking

  - Be sure to remove plate sealer before putting in reader!

  - You will need to turn reader on at least 40 minutes before you are ready to read plate for warm-up and calibration.  The end of the 1st incubation with beads is a good time to start it up.

- Cell Lysis buffer

  - Preliminary tests with the phospho kits showed that the DAL lab lysis buffer recipe used for the kinase assays does not work.  Similar results were seen for p-IB, p-p38, p-ERK1/2, and p-Akt.  Although not always as severe, DAL lab lysis buffer generally had lower signal.

  - Cell lysis buffer is made up using the recipe below.  In addition to the amount you need for lysing (I’ve used the same levels & protocol as I did with DAL lab lysis buffer), you need some for the BCA Assay (usually about 500uL) and then for diluting for running on the plate (about 50-55uL per sample if a concentrated lysate).  You can always make more up so don’t waste it!!

| 990 uL | Lysis buffer (in cold room) |
| --- | --- |
| 4 uL | Factor A (in -20 space) |
| 2 uL | Factor B (in -20 space) |
| 4 uL | PMSF (in -20 space) – add immediately before lysing |
| the frozen components should be thawed on ice, and are good up to 5 thaws (please mark on tube). If you are on thaw 4 and there is a lot left, please make aliquots to help make it last! |

- Rereading a plate

  - In general, can read 2-3 times and over a 1 month period if kept in dark and at 4oC

  - To reread, vacuum filter plate and resuspend as you did at the end of protocol

- Phosphoprotein specific

  - Kits come with controls, which are worth running the first time.  They tell you that the kit worked and you did it correctly

  - After that, if you have lysates that will give signals on the full multiplex they may not be worthwhile

  - There is a fairly high assay variation (10-20%).  May be worth running a master lysate which has medium signal for all parts of the multiplex for normalization.  This could also serve as a replacement for the controls.

  - Need to run a lysis buffer blank

  - Linearity…  For some antibodies see a linear response to protein level (data here for HT-29’s, note protein level is ug/well not ug/mL – this trend seen for p-Akt, p-JNK, p-IKBa, p-p38 MAPK) while for other’s this trend is not there (p-ERK1/2).  However, for p-ERK1/2 there is a trend for phosphorylation signal which is linear if you blend a high and low signal lysate, indicating the the ratios are valid.

- Cytokine

  - Sample and standards should be diluted in your tissue culture media – ideally from the same bottle as used for the experiment for minimal variation

  - Due to number of wells you need to devote for standard curve, probably best to only split a plate for 2 assays at most